## Summary of SKD Kits Pendings
Model | BLT Invoice No. | Received Date | Received Qty | Problem QTY | Problem reported date to BLT | Date | No of days delay
Tsruscope III- Basic | 24ZXS121903 | 2024-12-31 | 500 | 77 | 2025-02-12 | 2025-05-14 | 91
| 25ZXS030401 | 2025-03-15 | 500 | 6 | 2025-04-07 | 2025-05-14 | 37
Truscope III- Analog | 25ZXS022001 | 2025-03-05 | 250 | 9 | 2025-03-17 | 2025-05-14 | 58
| 25ZXS041102 | 2025-04-25 | 100 | 1 | 2025-04-30 | 2025-05-14 | 14
Truscope M10 with Touch | 24ZXS101501 | 2024-10-26 | 50 | 28 | 2025-01-11 | 2025-05-14 | 123
Truscope M10 | 24ZXS112101 | 28-12-2024 | 500 | 85 | 2025-01-11 | 2025-05-14 | 123
Truscope Ultra Q5 - Digital | 25ZXS020901 | 2025-02-25 | 100 | 15 | 2025-03-17 | 2025-05-14 | 58
Truscope Ultra Q5 - Nellcor | 25ZXS041402 | 2025-04-25 | 50 | 5 | 2025-04-30 | 2025-05-14 | 14
Truscope III -Nellcor | 25ZXS031103 | 2025-03-25 | 250 | 6 | 2025-04-26 | 2025-05-14 | 18
Total |  |  | 2300 | 232

## PIVOT
Sum of Req QTY
Model2 | Spares required | Total
TS-III | Battery | 83
| Back Cabinet | 11
| LCD | 10
| Battery lock | 8
| Front Cabinet | 7
| Silicon Keypad | 7
| Spo2 DOC10 Socket | 6
| Side plate sticker | 5
| SOP Board Power Source Sticker | 5
| Power source Metal frame | 5
| Backlight cable | 4
| IO Interface board | 4
| Spo2 extension cable | 3
| SMPS Module | 2
| Analog BM-100A module | 2
| LVDS Cable | 2
| Analog Module | 2
| Nellcor Spo2 I/F board | 2
| MB350 ARM board assembly | 1
| Power source frame | 1
Q5 | LCD | 12
| Arm Board | 5
| Double Channel slots | 4
| Keyboard Control Board | 4
| Rear connector plate | 3
| keyboard control board | 2
| BM100A module Digital with IBP | 1
| LVDS Cable | 1
M10 | LCD | 113
Grand Total |  | 315

## Summary Spares pendings
Model | Model | BLT Invoice No. | Received Date | Received Qty | Problem QTY | Problem reported date to BLT | Spares required | Req QTY | Date | No of days delay
Tsruscope III- Basic | TS-III | 24ZXS121903 | 2024-12-31 | 500 | 77 | 2025-02-12 | Battery | 77 | 2025-05-14 | 91
Tsruscope III- Basic | TS-III | [5 empty] | Battery lock | 2
Tsruscope III- Basic | TS-III | 25ZXS030401 | 2025-03-15 | 500 | 6 | 2025-04-07 | Back Cabinet | 6 | 2025-05-14 | 37
Tsruscope III- Basic | TS-III | [5 empty] | Front Cabinet | 2
Tsruscope III- Basic | TS-III | [5 empty] | LCD | 3
Tsruscope III- Basic | TS-III | [5 empty] | Backlight cable | 1
Tsruscope III- Basic | TS-III | [5 empty] | Power source Metal frame | 4
Tsruscope III- Basic | TS-III | [5 empty] | Battery | 1
Tsruscope III- Basic | TS-III | [5 empty] | Battery lock | 1
Tsruscope III- Basic | TS-III | [5 empty] | SOP Board Power Source Sticker | 1
Tsruscope III- Basic | TS-III | [5 empty] | Analog Module | 2
Tsruscope III- Basic | TS-III | [5 empty] | Spo2 DOC10 Socket | 3
Tsruscope III- Basic | TS-III | [5 empty] | Side plate sticker | 1
Tsruscope III- Basic | TS-III | [5 empty] | Analog BM-100A module | 1
Tsruscope III- Basic | TS-III | [5 empty] | SMPS Module | 1
Tsruscope III- Basic | TS-III | [5 empty] | IO Interface board | 3
Tsruscope III- Basic | TS-III | [5 empty] | Spo2 extension cable | 3
Truscope III- Analog | TS-III | 25ZXS022001 | 2025-03-05 | 250 | 9 | 2025-03-17 | Spo2 DOC10 socket | 2 | 2025-05-14 | 58
Truscope III- Analog | TS-III | [5 empty] | Analog BM-100A module | 1
Truscope III- Analog | TS-III | [5 empty] | MB350 ARM board assembly | 1
Truscope III- Analog | TS-III | [5 empty] | Battery | 5
Truscope III- Analog | TS-III | [5 empty] | SOP Board Power Source Sticker | 4
Truscope III- Analog | TS-III | [5 empty] | Battery Lock | 1
Truscope III- Analog | TS-III | [5 empty] | Power source Metal frame | 1
Truscope III- Analog | TS-III | [5 empty] | LCD | 5
Truscope III- Analog | TS-III | [5 empty] | LVDS Cable | 2
Truscope III- Analog | TS-III | [5 empty] | Backlight Cable | 2
Truscope III- Analog | TS-III | [5 empty] | Silicon Keypad | 4
Truscope III- Analog | TS-III | [5 empty] | Front Cabinet | 2
Truscope III- Analog | TS-III | [5 empty] | Back Cabinet | 3
Truscope III- Analog | TS-III | 25ZXS041102 | 2025-04-25 | 100 | 1 | 2025-04-30 | LCD | 1 | 2025-05-14 | 14
Truscope III- Analog | TS-III | [5 empty] | Silicon Keypad | 1
Truscope III- Analog | TS-III | [5 empty] | Back Cabinet | 1
Truscope III -Nellcor | TS-III | 25ZXS031103 | 2025-03-25 | 250 | 6 | 2025-04-26 | Back Cabinet | 1 | 2025-05-14 | 18
Truscope III -Nellcor | TS-III | [5 empty] | Front Cabinet | 3
Truscope III -Nellcor | TS-III | [5 empty] | Silicon Keypad | 2
Truscope III -Nellcor | TS-III | [5 empty] | LCD | 1
Truscope III -Nellcor | TS-III | [5 empty] | Backlight cable | 1
Truscope III -Nellcor | TS-III | [5 empty] | Battery lock | 4
Truscope III -Nellcor | TS-III | [5 empty] | Power source frame | 1
Truscope III -Nellcor | TS-III | [5 empty] | Side plate sticker | 4
Truscope III -Nellcor | TS-III | [5 empty] | Nellcor Spo2 I/F board | 2
Truscope III -Nellcor | TS-III | [5 empty] | Spo2 DOC10 socket | 1
Truscope III -Nellcor | TS-III | [5 empty] | SMPS Module | 1
Truscope III -Nellcor | TS-III | [5 empty] | IO Interface board | 1
Truscope M10 with Touch | M10 | 24ZXS101501 | 2024-10-26 | 50 | 28 | 2025-01-11 | LCD | 28 | 2025-05-14 | 123
Truscope M10 | M10 | 24ZXS112101 | 28-12-2024 | 500 | 85 | 2025-01-11 | LCD | 85 | 2025-05-14 | 123
Truscope Ultra Q5 - Nellcor | Q5 | 25ZXS041402 | 2025-04-25 | 50 | 5 | 2025-04-30 | Arm Board | 2 | 2025-05-14 | 14
Truscope Ultra Q5 - Nellcor | Q5 | [5 empty] | LVDS Cable | 1
Truscope Ultra Q5 - Nellcor | Q5 | [5 empty] | keyboard control board | 2
Truscope Ultra Q5 - Nellcor | Q5 | [5 empty] | LCD | 3
Truscope Ultra Q5 - Digital | Q5 | 25ZXS020901 | 2025-02-25 | 100 | 15 | 2025-03-17 | Keyboard Control Board | 4 | 2025-05-14 | 58
Truscope Ultra Q5 - Digital | Q5 | [5 empty] | Double Channel slots | 4
Truscope Ultra Q5 - Digital | Q5 | [5 empty] | BM100A module Digital with IBP | 1
Truscope Ultra Q5 - Digital | Q5 | [5 empty] | Rear connector plate | 3
Truscope Ultra Q5 - Digital | Q5 | [5 empty] | Arm Board | 3
Truscope Ultra Q5 - Digital | Q5 | [5 empty] | LCD | 9
Total |  |  |  | 2300 | 232 |  |  | 315

## TS-III overall Report (hidden)
Model | Schiller Received Date | BLT Invoice No. | Received Qty |  | Defective Qty in Production
TSIII- Baisc Qty- 500 | 2024-08-30 | 24ZXS080101-1 & 24ZXS080101-2 | 500 | 22 | 0
TSIII - Upgradable Qty-50 | 2024-08-30 | 24ZXS080102-1 & 24ZXS080102-3 | 50 | 1 | 0
TSIII -Nellcor Upgradable Qty-100 | 2024-09-19 | 24ZXS090302-1 & 24ZXS090302-3 | 98 | 2 | 0
Truscope - M10 Qty-450 (435 + 15 VGA) | 24-09-2024 10-10-2024 | 24ZXS090301-2 & 24ZXS090301-1 | 443 | 3 | 4
Truscope III- upgradable Qty-100 (Analog) | 22-10-2024 19-10-2024 | 24ZXS100901-1 & 24ZXS100901-3 | 97 | 3 | 0
Truscope III- upgradable Qty-100 (Nellcor) | 2024-10-19 | 24ZXS100901-3 | 99 | 0 | 1
Truscope - M10 Qty-50 Nos (with Touch) | 2024-10-26 | 24ZXS101501-1 & 24ZXS101501-2 | 49 | 1 | 0
TSIII- Baisc Qty-500 (Non upgradable) | 2024-10-30 | 24ZXS092401-1 & 24ZXS092401-3 | 476 | 24 | 0
Truscope Ultra Q5 Qty-50 Nos (Digital) | 2024-11-06 | 24ZX102101-3 & 24ZX102101-1 | 36 | 14 | 0
Truscope III- upgradable Qty-100 (Analog) | 2024-11-12 | 24ZXS110505-1 & 24ZXS110505-3 | 99 | 1 | 0
|  |  | 98 | 2 | 0
TS III -Nellcor Qty-50 Nos | 2024-11-28 | 24ZXS111502-3 | 50 | 0 | 0
Ultra Q5 Received Qty-100 Nos | 2025-01-10 | 24ZXS123001-3 & 24ZXS123001-1 | 82 | 17 | 1
TS III -Nellcor Upgradable Qty-200 Nos | 10-01-2025 18-01-2025 | 24ZXS123101-3 & 24ZXS123101-1 | 196 | 2 | 2
TS III - Masimo Qty - 15 | 2024-12-10 | 24ZXS123101-3 | 12 | 0 | 3
| 2025-01-18 | 24ZXS123101-1 | 13 | 2 | 0
| 2025-01-20 | 24ZXS123101-2 | 14 | 1 | 0
Truscope III Qty-500 (485 +15) with VGA | 2025-02-04 | 24ZXS121903-1 24ZXS121903-3 |  | 31 | 3
Truscope M10 Qty - 500 (460+25+15) | 28-12-2024 | 24ZXS112101-2 24ZXS112101-1 | 494 | 6 | 0

## TS-III Basic SKD Issues
Received Date | BLT Invoice No. | Received Qty | Total SKD kits cannot be make saleable | P/N | Defective Item | Spares required. | Failure Qty | Rejection folder
2024-12-31 | 24ZXS121903 | 500 | 77 | 13.0005 | Battery | Rust on battery terminal and we require battery -77 No's Battery lock - 2 No's | 77
2025-03-15 | 25ZXS030401 | 500 | 6 | 60.000850A | Back cabinet(cabinet assembly) | Back Cabinet -6 nos | 6 | https://drive.google.com/drive/folders/1pg-jIfE0mhG1j6UMY4w6Y3Xs1FCXbMDi
|  |  |  | 60.000840A | Front cabinet(cabinet assembly) | Front Cabinet - 2 nos | 2
|  |  |  | 60.000802A | LCD assembly | LCD -3 nos. Backlight cable - 1 No | 3
|  |  |  | 60.000851A | Power source assembly frame | Frame - 4 nos Battery - 1no Battery lock-1no power source sticker -1no | 5
|  |  |  | 60.000858A | Side Plate Assy(2-IBP 1-TEMP with Analog SpO2) | Analog Module - 2 Spo2 DOC10 Socket - 3 Side plate sticker - 1no | 6
|  |  |  | 60.000856A | Digital BM-100A module with SpO2 | BM100A Module - 1 no | 1
|  |  |  | 50.000302A | SMPS module | SMPS Module-1 | 1
|  |  |  | 60.000743A | MB350 arm board assembly | IO Board -3 Nos. | 3
|  |  |  | 15-100-0349 | Analog spo2 extension cable | Spo2 extension cable - 3 nos | 3

## TS-III Analog SKD issues
Received Date | BLT Invoice No. | Received Qty | Total SKD kits cannot be make saleable | P/N | Defective Item | Description | Failure Qty | Spares required. | Serial No./Lot No. | Rejection Images
2025-03-05 | 25ZXS022001 | 250 | 9 | 60.000859A | Side Plate Assembly(2-IBP; 2-TEMP,analog spo2) | SpO2 Waveform not coming - 2 Nos | 2 | Spo2 DOC10 socket -2 Nos | 09ZD0004403, 09ZD0004830 | https://drive.google.com/drive/folders/1jMYEWMX8A0rotngfutG1f6U1s8sJxK_m
|  |  |  | 60.000857A | Digital BM-100A module | SpO2 Base line not Coming - 1No. | 1 | Analog BM-100A module -1 No | 00ZE0000245
|  |  |  | 60.000843A | MB350 arm board assembly | Unit Not Switching On-1No | 1 | MB350 ARM board assembly - 1 No | AJ-D24110501
|  |  |  | 60.000841A | Power source assembly frame | Sticker Scratch and Battery terminal Rust - 1No. Battery terminal Rust - 3. Power socket bend and Battery terminal Rust -1 No. Sticker Scratch- 2 No. Sticker Scratch and Battery Lock Damage - 1No. | 8 | Battery- 5 Nos. SOP Board Power Source Sticker - 4 Nos. Battery Lock - 1 No. Power source Metal frame- 1No
|  |  |  | 60.000802A | LCD assembly | ⁠Dust on LCD -1 No ⁠Dot issue-4 nos ( Green dot -1 no.Blue Dot-3). Cable Damage-2nos ( Backlight Cable Damage - 2 nos). LVDS Connector Damage - 2 Nos | 9 | LCD - 5 Nos. Backlight Cable - 2 Nos. LVDS Cable - 2 Nos. | NL-3U25010495, NL -3U25010883, NL-3U25010848, NL -3U25010494, NL -3U25010625, NL-3U25010562, NL-3U25010929, NL-3U25010656, NL-3U25010691
|  |  |  | 60.000840A | Front cabinet(cabinet assembly) | Sillicon Keypad peeloff - 3 Nos. Black Dot on Cabinet - 1Nos. Printing problem in Sillicon Keypad -1 Printing issue in cabinet - 1 Nos. | 6 | Sillicon Key Pad - 4 Nos. Cabinet - 2 Nos.
|  |  |  | 60.000854A | Back cabinet(cabinet assembly) | Inserrt Missing - 1 No. Black Dot on Cabinet - 1No. Cabinet Lock Damage - 1 No. | 3 | Cabinet - 3 Nos.
25-04-2025 | 25ZXS041102 | 100 | 1 | 60.000802A | LCD assembly | Blue dot on LCD -1 No. | 1 | LCD- 1 No | NL-3U24101598 | https://drive.google.com/drive/folders/11SzOXtiEJ6XUg1wt5mj9UKHBoKGAdWAo
|  |  |  | 60.000840A | Front cabinet(cabinet assembly) | Dust on sillicon keypad - 1 No | 1 | Sillicon Keypad - 1 No
|  |  |  | 60.000854A | Back cabinet(cabinet assembly) | Printing issue in Cabinet - 1 No | 1 | Back Cabinet - 1 No

## TS-III Nellcor SKD issues
Received Date | BLT Invoice No. | Received Qty | Total SKD kits cannot be make saleable | P/N | Defective Item | Spares required QTY | Failure Qty | Serial No./Lot No. | Rejection Images / Folder
2025-03-25 | 25ZXS031103 | 250 | 6 | 60.000850A | Back cabinet(cabinet assembly) | Back Cabinet -1 No. | 1 |  | https://drive.google.com/drive/folders/1tBpl7vX40VPHR7phk7J9pX6wQLj21FQ6
|  |  |  | 60.000840A | Front cabinet(cabinet assembly) | Front Cabinet - 3Nos. Silicon keypad - 2Nos. | 5
|  |  |  | 60.000802A | LCD assembly | LCD -1 No. Backlight cable - 1 No | 1 | NL-3U24120316
|  |  |  | 60.000841V | Power source assembly frame with VGA-TS III | Battery lock- 4 Nos. Power source frame - 1no | 5
|  |  |  | 60.000844A | Side Plate Assembly(2-IBP; 2-TEMP,Nellcor) | Side plate sticker - 4 Nos. Nellcor Spo2 I/F board - 2no DOC-10 connector - 1no | 6
|  |  |  | 50.000302A | SMPS | SMPS Module-1 No. | 1 | XH-1N25020133
|  |  |  | 60.000843V | MB350 arm board assembly with VGA-TS III | I/O Interface board - 1no | 1 | 17ME000101

## Q5 Digital SKD issues
Received Date | BLT Invoice No & Date | Received Qty | Total SKD kits cannot be make saleable | P/N | Part description | Problem Description | Failure QTY | Spares required QTY | Serial No./Lot No. | Reference images | Final Remarks
2025-02-25 | 25ZXS020901 | 100 | 15 | 60.0005 | LCD assembly | 1. J-12 Connector damage - 3 Nos. 2. J-4 Connector damage - 1No. 3. Blue Line on LCD - 9 No's 4. Flickering Issue - 2 No's | 15 | 1. Keyboard Control Board - 4 Nos. 2. LCD - 9 Nos | NL-3U24120711, NL-3U24121156 NL-3U24120678, NL-3U24120773 NL-3U24120869, NL-3U24120753, NL-3U24120707,NL-3U24120784, NL-3U24120706,NL-3U24120750, NL-3U24101066,NL-3U24120780, NL-3U24121154,NL-3U24121083, NL-3U24121154 | https://drive.google.com/drive/folders/1kRoaZ3SfxQfIV2wPlYUw-e8snUHIafVV | Not Dispatched
|  |  |  | 60.0005 | Double Channel slots | 1. Connector Damage - 3 nos. 2. U1 Processor burnt - 1 No. | 4 | Double Channel slots - 4 no's
|  |  |  | 60.0005 | BM100A Module with IBP - Digital | ECG Communication Error | 1 | BM100A module Digital with IBP - 1 no | 00ZE0000363
|  |  |  | 60.0005 | Rear connector | J-3 Connector Damege - 3Nos | 3 | Rear connector plate - 3 no's
|  |  |  | 60.0005 | Arm board assembly with power source | 1. Flickering Issue -2 No. 2. Unit Not Switching ON -1No | 3 | Arm Board - 3 no's | AJ-G24100445, AJ-G24010196 AJ-G24100819

## Q5 Nellcor SKD issues
Received Date | BLT Invoice No. | Received Qty | Total SKD kits cannot be make saleable | P/N | Defective Item | Problem Description | Spares required QTY | Serial No./Lot No. | Rejection images/Folder
25/04/2025 | 25ZXS041402 | 50 | 5 | 60.0005 | Arm board assembly with power source | Sequence number missing - 2nos | Arm Board - 2nos. | XH-1W24120013, XH-1W24120005 | https://drive.google.com/drive/folders/11m6kFc9kumNk1fr2tEAFELfp1RHi9_H2
|  |  |  | 60.0005 | LCD assembly | 1. Dot issues (Black dot) - 2nos. 2. Green dot + Keyboard Control Board connector damage - 1 No 3. Keyboard Control Board Connector Damage - 1 no. 4. LVDS Cable damage - 1no. | 1. LCD - 3 No's. 2. keyboard control board - 2 No's. 3. LVDS Cable - 1 No. | NL-3U24100719,NL-3U24100805, NL-3U24100796,NL-3U24101266, NL-3U24100793

## M10 SKD issues
Received Date | BLT Invoice No. | Received Qty | Total SKD kits cannot be make saleable | P/N | Defective Item | Spares required QTY | Failure Qty
28-12-2024 | 24ZXS112101 | 500 | 85 | 1.14.101-0010-00-1 | 10.4' LCD | 1. Dull display - 74No's 2. Dot on the display.(Red, Green, Black) - 9No's 3. Rust on insert - 1No 4. Vertical line - 1No | 85
2024-10-26 | 24ZXS101501 | 50 | 28 | 1.14.101-0010-00-1 | 10.4' LCD | Brightness issue | 28

## Replacement Rejection
Model | Item No. | Description | Failure details | Qty | Import invoice & Date details | Final Remarks
TSIII -Upgradable | 50.0003 | Battery lock of TS III | Wrong item recv'd | 4 | 24HGZ120601 dt. 06-12-2024 | Complete battery lock not received
Mini | 60.0004 | DIGITAL SPO2 MODULE | Sensor not detected - Z20448737 | 1 | 24HGZ121102 dt.11.12.2024 | Received NIBP Module instead of Spo2 module
TSIII - Upgradable | 50.0003 | Battery lock of TS III | Wrong item Rece'd | 5 | 24HGZ121102 dt.11.12.2024 | Complete battery lock not received
TSIII - Upgradable | 60.000859A | Side Plate Assembly(2-IBP; 2-TEMP,analog spo2) | Side connector panel problem, I/F board & SPO2 Module and sticker scratches | 1 | 24HGZ121102 dt.11.12.2024 | Analog Interface board and analog module not received. Only sticker received.
M10 | 1.14.101-0010-00-1 | 10.4' LCD | LCD Dull Display | 6 | 24HGZ121102 dt.11.12.2024 | Received LCD is dull display. Need replacement
M10 | 1.14.101-0010-00-1 | 10.4' LCD | LCD Dull Display | 7 | 24HGZ121101 dt.11.12.2024 | Received LCD is dull display. Need replacement
TS-III Masimo | 60.000845A | Side Plate Assembly(2-IBP; 2-TEMP,masimo) | Masimo module problem(Spo2 wave form not coming) and also Sticker scratches-3nos | 3 | 25HGZ040301 dt.03.04.2025 | Masimo module 2 no's not received instead of that, we received only the side plate