## Summary of SKD Kits Pendings
Model | BLT Invoice No. | Received Date | Received Qty | Problem QTY | Problem reported date to BLT | Replacement Status
Tsruscope III- Basic | 24ZXS080101 | 2024-08-30 | 500 | 2 |  | Replacement Pending from BLT
| 24ZXS121903 | 2024-12-31 | 500 | 77 |  | Replacement Pending from BLT
| 24ZXS092401 | 2024-10-30 | 500 | 29 |  | Replacement Pending from BLT
| 25ZXS030401 | 2025-03-15 | 500 | 6 | 2025-04-07 | Replacement Pending from BLT
Truscope III- Analog | 24ZXS080102 | 2024-08-30 | 50 | 1 |  | Replacement received against invoice 25HGZ040301 dt 03.04.2025
| 24ZXS110505 | 2024-11-12 | 100 | 27 |  | Replacement received against invoice 25HGZ040301 dt 03.04.2025
| 25ZXS022001 | 2025-03-05 | 250 | 9 | 2025-03-17 | Replacement Pending from BLT
Truscope III -Nellcor | 24ZXS123101 | 2025-01-10 | 200 | 10 |  | Replacement Pending from BLT
Truscope III - Masimo | 24ZXS123101 | 2024-12-10 | 15 | 3 |  | Replacement received against invoice 25HGZ040301 dt 03.04.2025
Truscope M10 with Touch | 24ZXS101501 | 2024-10-26 | 50 | 28 |  | BLT planning to dispatch the replaced spares (Except LCD) against invoice 25HGZ041501 on 23rd April'25
Truscope M10 | 24ZXS112101 | 28-12-2024 | 500 | 85 |  | BLT planning to dispatch the replaced spares (Except LCD) against invoice 25HGZ041501 on 23rd April'25
Truscope Ultra Q5 - Digital | 25ZXS020901 | 2025-02-25 | 100 | 15 | 2025-03-17 | Replacement Pending from BLT
| 24ZXS102101 | 2024-11-06 | 50 | 2 |  | BLT planning to dispatch the replaced spares against invoice 25HGZ041501 on 23rd April'25
Truscope Ultra Q5 - Nellcor | 24ZXS102101 | 2024-11-06 | 50 | 1 |  | BLT planning to dispatch the replaced spares against invoice 25HGZ041501 on 23rd April'25
| 24ZXS123001 | 2025-01-10 | 100 | 20 |  | BLT planning to dispatch the replaced spares against invoice 25HGZ041501 on 23rd April'25
Truscope III -Nellcor | 25ZXS031103 | 2025-03-25 | 250 | 6 | 2025-04-26 | Replacement Pending from BLT
Total |  |  | 3715 | 321

## TS-III overall Report (hidden)
Model | Schiller Received Date | BLT Invoice No. | Received Qty |  | Defective Qty in Production
TSIII- Baisc Qty- 500 | 2024-08-30 | 24ZXS080101-1 & 24ZXS080101-2 | 500 | 22 | 0
TSIII - Upgradable Qty-50 | 2024-08-30 | 24ZXS080102-1 & 24ZXS080102-3 | 50 | 1 | 0
TSIII -Nellcor Upgradable Qty-100 | 2024-09-19 | 24ZXS090302-1 & 24ZXS090302-3 | 98 | 2 | 0
Truscope - M10 Qty-450 (435 + 15 VGA) | 24-09-2024 10-10-2024 | 24ZXS090301-2 & 24ZXS090301-1 | 443 | 3 | 4
Truscope III- upgradable Qty-100 (Analog) | 22-10-2024 19-10-2024 | 24ZXS100901-1 & 24ZXS100901-3 | 97 | 3 | 0
Truscope III- upgradable Qty-100 (Nellcor) | 2024-10-19 | 24ZXS100901-3 | 99 | 0 | 1
Truscope - M10 Qty-50 Nos (with Touch) | 2024-10-26 | 24ZXS101501-1 & 24ZXS101501-2 | 49 | 1 | 0
TSIII- Baisc Qty-500 (Non upgradable) | 2024-10-30 | 24ZXS092401-1 & 24ZXS092401-3 | 476 | 24 | 0
Truscope Ultra Q5 Qty-50 Nos (Digital) | 2024-11-06 | 24ZX102101-3 & 24ZX102101-1 | 36 | 14 | 0
Truscope III- upgradable Qty-100 (Analog) | 2024-11-12 | 24ZXS110505-1 & 24ZXS110505-3 | 99 | 1 | 0
|  |  | 98 | 2 | 0
TS III -Nellcor Qty-50 Nos | 2024-11-28 | 24ZXS111502-3 | 50 | 0 | 0
Ultra Q5 Received Qty-100 Nos | 2025-01-10 | 24ZXS123001-3 & 24ZXS123001-1 | 82 | 17 | 1
TS III -Nellcor Upgradable Qty-200 Nos | 10-01-2025 18-01-2025 | 24ZXS123101-3 & 24ZXS123101-1 | 196 | 2 | 2
TS III - Masimo Qty - 15 | 2024-12-10 | 24ZXS123101-3 | 12 | 0 | 3
| 2025-01-18 | 24ZXS123101-1 | 13 | 2 | 0
| 2025-01-20 | 24ZXS123101-2 | 14 | 1 | 0
Truscope III Qty-500 (485 +15) with VGA | 2025-02-04 | 24ZXS121903-1 24ZXS121903-3 |  | 31 | 3
Truscope M10 Qty - 500 (460+25+15) | 28-12-2024 | 24ZXS112101-2 24ZXS112101-1 | 494 | 6 | 0

## TS-III Basic SKD Issues
Received Date | BLT Invoice No. | Received Qty | Total SKD kits cannot be make saleable | P/N | Defective Item | Description | Failure Qty
2024-08-30 | 24ZXS080101 | 500 | 2 | 60.000743A | MB350 arm board assembly-TS III lite | 1. Arm board-IC burnt, I/O board-Not getting ON-1 No. 2. Arm board-Display Hang, I/O board-Battery % not showing-1 No. | 2
2024-10-30 | 24ZXS092401 | 500 | 29 | 60.0008 | BACK PANEL STICKER POWER SOURCE ASSY TS III BASIC | Power source assembly frame-TS III back panel sticker scratches-4 nos , Spring damaged-1 no. | 4
|  |  |  | 50.0003 | Battery lock of TS III | Power source assembly frame-TS III Battery lock damaged | 4
|  |  |  | 60.000743A | MB350 arm board assembly-TS III lite | Key board function not working. SN: AJ-D24070186 | 1
|  |  |  | 60.000856A | Digital BM-100A module with BM20 Analog sp2-TS III | ECG noise. SN: 00ZD00010820 | 1
|  |  |  | 175.0001 | Fan for TS III | Fan noise | 22
|  |  |  | 60.000820 | I/O BOARD | Unit not getting on-2 Nos. SN: 17MD0006274 & 17MD0006057 | 2
|  |  |  | 60.000820 | I/O BOARD | Key function not working. SN: 17MD0006100 & 17MD0006308 & 17MD0005518 | 3
|  |  |  | 60.000820 | I/O BOARD | Battery % not shows full after full charge. SN: 17MD0006195 | 1
|  |  |  | 50.000260 | LVDS cable | LVDS cable damage from board end | 1
|  |  |  | 50.000230 | Spo2 IF board | Spo2 base line not coming | 2
|  |  |  | 60.0008 | Speaker | Speaker not working | 1
|  |  |  | 2.00.0048-03-01 | Spo2 Module | Spo2 waveform not coming. SN: 09ZD0002667 & 09ZD0003347 & 09ZD0003046 | 3
|  |  |  | 2.00.0048-03-01 | Spo2 Module | Spo2 base line not coming. SN: 09ZD0002574 | 1
|  |  |  | 60.000851A | Power source assembly frame | Thread not there for Arm board fixing spacer | 1
|  |  |  | 60.000851A | Power source assembly frame | Power socket bend | 2
|  |  |  | 60.000851A | Power source assembly frame | Bottom side M4 spacer size mismatch | 1
|  |  |  | 60.000850A | Back cabinet assembly | Side plate fixing Insert is not there | 1
2024-12-31 | 24ZXS121903-01 | 500 | 77 | 60.000840A | Front cabinet(cabinet assembly) | 1. Cabinet scartch and silicon key printing peeloff -13 No's 2. Silicon key printing not proper -11No's 3. Cabinet insert missing and Silicon key printing peeloff - 1No's 4. Silicon key printing not proper and peel off - 3No's 5. Silicon key printing peeloff - 3No's, Cabinet with keyboard - 3 no's | 34
|  |  |  | 60.000850A | Back cabinet assembly | 1. Black dot - 5 Nos 2. Printing issue - 1No. 3. Scratch -1 No., Fan noise - 4nos Speaker sound not come - 3nos Rear cabinet handle scratch - 1no | 11
|  |  |  | 60.000802A | LCD assembly | Low brightness-1No., Dot issue-5 Nos ( 3 nos balck dot, 1 nos green dot , 1 nos red dot) SN. NL-3U24110486,NL-3U24110713, NL-3U24110015, Frame damage + green dot - 1 No, Frame damage + corner dot - 1 No, SN. NL-3U24110616, NL-3U24110494 NL-3U24110824, NL-3U24110734 NL-3U24110901, backlight cable damage-15 nos, LVDS cable damaged, 2 nos | 21
|  |  |  | 13.0005 | Battery | Rust on battery terminal-77 No's, Including (Battery lock Damage- 2 No's, sticker scratch- 6 No's) | 77
|  |  |  | 60.000858A | Side Plate Assy(2-IBP 1-TEMP with BM20 Analog sp2) | sticker scratch, Spo2 connector one side not fixed properly - 5nos Spo2 wave form not come - 2nos Spo2 sensor not detected - 1no | 6
|  |  |  | 60.000743A | MB350 arm board assembly-TS III lite | Unit not Switching ON, Not getting on - 4nos (IF board) SN.17MD0006718, 17MD0007038,17MD0006795,17MD0006675,17MD0006687, 17MD0003254 AJ-D24100928,AJ-D24111149,AJ-E23030017 Alarm (Orange) LED not glowing - 1no (IF Board) Unit hanging & key function not working - 3nos (ARM Board) | 6
|  |  |  | 60.000856A | Digital BM-100A module with BM20 Analog sp2-TS III | ECG Noise - 1no SPO2 Base line not come - 1no NIBP Not working (intermidiate) - 1no, SN.00ZD0012487 00ZD0013972 00ZD0014075 | 3
2025-03-15 | 25ZXS030401 | 500 | 6 | 60.000850A | Back cabinet(cabinet assembly) | Back Cabinet -6 nos | 6 | https://drive.google.com/drive/folders/1pg-jIfE0mhG1j6UMY4w6Y3Xs1FCXbMDi
|  |  |  | 60.000840A | Front cabinet(cabinet assembly) | Front Cabinet - 2 nos | 2
|  |  |  | 60.000802A | LCD assembly | LCD -3 nos. Backlight cable - 1 No | 3
|  |  |  | 60.000851A | Power source assembly frame | Frame - 4 nos Battery - 1no Battery lock-1no power source sticker -1no | 5
|  |  |  | 60.000858A | Side Plate Assy(2-IBP 1-TEMP with Analog SpO2) | Analog Module - 2 Spo2 DOC10 Socket - 3 Side plate sticker - 1no | 6
|  |  |  | 60.000856A | Digital BM-100A module with SpO2 | BM100A Module - 1 no | 1
|  |  |  | 50.000302A | SMPS module | SMPS Module-1 | 1
|  |  |  | 60.000743A | MB350 arm board assembly | IO Board -3 Nos. | 3
|  |  |  | 15-100-0349 | Analog spo2 extension cable | Spo2 extension cable - 3 nos | 3

## TS-III Analog SKD issues
Received Date | BLT Invoice No. | Received Qty | Total SKD kits cannot be make saleable | P/N | Item Name | Description | Failure Qty | Additional Remarks | Serial No./Lot No. | Rejection Images | Final Remarks
2024-08-30 | 24ZXS080102 | 50 | 1 | 60.000843A | MB350 arm board assembly-TS III | Arm board-Blank display-1 | 1 |  |  |  | Given from service
2024-11-12 | 24ZXS110505 | 100 | 7 | 60.000840A | Front cabinet(cabinet assembly) | Cabinet scratches, alarm lamp cover missing | 27 |  |  |  | Received
|  |  |  | 60.000854A | Back cabinet(cabinet assembly)-TS III | Symbol reprinted-1 No., Cabinet mould finishing problem-1 No. | 2 |  |  |  | Received
|  |  |  | 60.0008 | Power source assembly frame with VGA-TS III | Sticker scratches-7 Nos | 7 |  |  |  | Received
|  |  |  | 60.000843V | MB350 arm board assembly with VGA-TS III | Not getting ON | 1 |  |  |  | Given from service
|  |  |  | 60.000841V | Power source assembly frame with VGA-TS III | Connector base plate bend-1 No. Frame bend near power socket-1 No. | 2 |  |  |  | Received
|  |  |  | 60.000859A | Side Plate Assembly(2-IBP; 2-TEMP,analog spo2) | Side plate scratches | 1 |  |  |  | Received
|  |  |  | 60.000820 | I/O BOARD | Not getting on. SN: 17MD0006926, 17MD0006934, 17MD0006867, 17MD0006853, 17MD0006846, 17MD0006924 | 6 |  |  |  | Received
|  |  |  | 60.000820 | I/O BOARD | Key function not working & mains symbol not showing. SN: 17MD0006945 | 1 |  |  |  | Received
|  |  |  | 2.00.0048-03-01 | Spo2 Module | Spo2 waveform not coming. SN:09ZD0003032, 09ZD0003098 | 2 |  |  |  | Received
|  |  |  | 2.00.0048-03-01 | Spo2 Module | Spo2 base line not coming. SN: 09ZD0003035 | 1 |  |  |  | Received module not working
|  |  |  | 60.000857A | Digital BM-100A module | SP02 IF baord fixing spacer missing | 1 |  |  |  | Received
|  |  |  | 60.000803 | SILICA GEL KEYBOARD | Alarm disable symbol is not proper | 1 |  |  |  | Received
|  |  |  | 60.0008 | 12.1' LCD | Display color changed. SN: NL-3U24050068 | 1 |  |  |  | Received
2025-03-05 | 25ZXS022001 | 250 | 9 | 60.000859A | Side Plate Assembly(2-IBP; 2-TEMP,analog spo2) | SpO2 Waveform not coming - 2 Nos | 2 | Spo2 DOC10 socket -2 Nos | 09ZD0004403, 09ZD0004830 | https://drive.google.com/drive/folders/1jMYEWMX8A0rotngfutG1f6U1s8sJxK_m
|  |  |  | 60.000857A | Digital BM-100A module | SpO2 Base line not Coming - 1No. | 1 | Digital BM-100A module -1 No | 00ZE0000245
|  |  |  | 60.000843A | MB350 arm board assembly | Unit Not Switching On-1No | 1 | MB350 ARM board assembly - 1 No | AJ-D24110501
|  |  |  | 60.000841A | Power source assembly frame | Sticker Scratch and Battery terminal Rust - 1No. Battery terminal Rust - 3. Power socket bend and Battery terminal Rust -1 No. Sticker Scratch- 2 No. Sticker Scratch and Battery Lock Damage - 1No. | 8 | Battery- 5 Nos. SOP Board Power Source Sticker - 4 Nos. Battery Lock - 1 No. Power source Metal frame- 1No
|  |  |  | 60.000802A | LCD assembly | ⁠Dust on LCD -1 No ⁠Dot issue-4 nos ( Green dot -1 no.Blue Dot-3). Cable Damage-2nos ( Backlight Cable Damage - 2 nos). LVDS Connector Damage - 2 Nos | 9 | LCD - 5 Nos. Backlight Cable - 2 Nos. LVDS Cable - 2 Nos. | NL-3U25010495, NL -3U25010883, NL-3U25010848, NL -3U25010494, NL -3U25010625, NL-3U25010562, NL-3U25010929, NL-3U25010656, NL-3U25010691
|  |  |  | 60.000840A | Front cabinet(cabinet assembly) | Sillicon Keypad peeloff - 3 Nos. Black Dot on Cabinet - 1Nos. Printing problem in Sillicon Keypad -1 Printing issue in cabinet - 1 Nos. | 6 | Sillicon Key Pad - 4 Nos. Cabinet - 2 Nos.
|  |  |  | 60.000854A | Back cabinet(cabinet assembly) | Inserrt Missing - 1 No. Black Dot on Cabinet - 1No. Cabinet Lock Damage - 1 No. | 3 | Cabinet - 3 Nos.

## TS-III Nellcor SKD issues
Received Date | BLT Invoice No. | Received Qty | Total SKD kits cannot be make saleable | P/N | Defective Item | Description | Failure Qty | Serial No./Lot No. | Rejection Images / Folder
2025-01-10 | 24ZXS123101 | 200 | 10 | 60.000802A | LCD assembly | 1. Red dots-2nos 2. Blank display-1 no., Horizandal line-1 no. | 4
|  |  |  | 60.000841V | Power source assembly frame with VGA-TS III | 1. Sticker scratches-5nos, 2. Battery lock damaged and also sticker scratches-2nos 3. Mains socket frame bend-1 no. | 8
|  |  |  | 60.000844A | Side Plate Assembly(2-IBP; 2-TEMP,Nellcor) | Sticker scratches-1no, Doc -10 Connector problem - 3no's | 4
|  |  |  | 60.000843V | MB350 arm board assembly with VGA-TS III | VGA FRC cable damaged-1no, I/O Board problem Unit not getting on - 1 no. | 2
|  |  |  | 60.000840A | Front cabinet(cabinet assembly) | Cabinet scratches-8nos, Encoder function not working properly-2 no's | 10
|  |  |  | 60.000854A | Back cabinet(cabinet assembly)-TS III | Black dots-6nos, speaker problem - 4no's | 10
2025-03-25 | 25ZXS031103 | 250 | 6 | 60.000850A | Back cabinet(cabinet assembly) | Back Cabinet -1 No. | 1 |  | https://drive.google.com/drive/folders/1tBpl7vX40VPHR7phk7J9pX6wQLj21FQ6
|  |  |  | 60.000840A | Front cabinet(cabinet assembly) | Front Cabinet - 3Nos. Silicon keypad - 2Nos. | 5
2025-03-27 |  |  |  | 60.000802A | LCD assembly | LCD -1 No. Backlight cable - 1 No | 1 | NL-3U24120316
|  |  |  | 60.000841V | Power source assembly frame with VGA-TS III | Battery lock- 4 Nos. Power source frame - 1no | 5
|  |  |  | 60.000844A | Side Plate Assembly(2-IBP; 2-TEMP,Nellcor) | Side plate sticker - 4 Nos. Nellcor Spo2 I/F board - 2no DOC-10 connector - 1no | 6
|  |  |  | 50.000302A | SMPS | SMPS Module-1 No. | 1 | XH-1N25020133
|  |  |  | 60.000843V | MB350 arm board assembly with VGA-TS III | Interface board - 1no | 1 | 17ME000101+L+I1:M8+I1:M8

## TS-III Masimo SKD issues
Received Date | BLT Invoice No. | Received Qty | Total SKD kits cannot be make saleable | P/N | Defective Item | Description | Failure Qty | Final Remarks
2024-12-10 | 24ZXS123101 | 15 | 3 | 60.000845A | Side Plate Assembly(2-IBP; 2-TEMP,masimo) | Masimo module problem(Spo2 wave form not coming) and also Sticker scratches-3nos | 3 | Masimo module 2 no's not received
|  |  |  | 4050 | Masimo adult spo2 sensor | Sensor LED not glowing and SpO2 waveform not coming | 1 | Received
|  |  |  | 60.000854A | Back Cabinet (Cabinet Assembly) | Black dot - 2 No's | 2 | Received

## Q5 Digital SKD issues
Received Date | BLT Invoice No & Date | Received Qty | Total SKD kits cannot be make saleable | P/N | Part description | Problem Description | Failure QTY | Additional remarks | Serial No./Lot No. | Reference images | Final Remarks
2024-11-06 | 24ZX102101 | 50 | 2 | 60.0005 | LCD Assembly | Blue veritical line-1 No. | 1 |  |  |  | Supplier planning to dispatch the replaced spares against invoice 25HGZ041501 on 23rd April'25
|  |  |  | 60.0005 | Arm Board | No Sequence number | 1
|  |  |  | 60.0009 | Back Cabinet | Cabinet top side damaged - 2 no's | 2
|  |  |  | 60.0005 | Side panel assembly | Temp2 not working | 1
|  |  |  | 60.0005 | Side key board | Key function not working properly | 1
2025-02-25 | 25ZXS020901 | 100 | 15 | 60.0005 | LCD assembly | 1. J-12 Connector damage - 3 Nos. 2. J-4 Connector damage - 1No. 3. Blue Line on LCD - 9 No's 4. Flickering Issue - 2 No's | 15 | 1. Keyboard Control Board - 4 Nos. 2. LCD - 9 Nos | NL-3U24120711, NL-3U24121156 NL-3U24120678, NL-3U24120773 NL-3U24120869, NL-3U24120753, NL-3U24120707,NL-3U24120784, NL-3U24120706,NL-3U24120750, NL-3U24101066,NL-3U24120780, NL-3U24121154,NL-3U24121083, NL-3U24121154 | https://drive.google.com/drive/folders/1kRoaZ3SfxQfIV2wPlYUw-e8snUHIafVV
|  |  |  | 60.0005 | Double Channel slots | 1. Connector Damage - 3 nos. 2. U1 Processor burnt - 1 No. | 4 | Double Channel slots - 4 no's
|  |  |  | 60.0005 | BM100A Module with IBP - Digital | ECG Communication Error | 1 | BM100A module Digital with IBP - 1 no | 00ZE0000363
|  |  |  | 60.0005 | Rear connector | J-3 Connector Damege - 3Nos | 3 | Rear connector plate - 3 no's
|  |  |  | 60.0005 | Arm board assembly with power source | 1. Flickering Issue -2 No. 2. Unit Not Switching ON -1No | 3 | Arm Board - 3 no's | AJ-G24100445, AJ-G24010196 AJ-G24100819

## Q5 Nellcor SKD issues
Received Date | BLT Invoice No. | Received Qty | Total SKD kits cannot be make saleable | P/N | Defective Item | Description | Failure Qty | Final Remarks
2025-01-10 | 24ZXS123001 | 100 | 20 | 60.0005 | LCD Assembly | Connector damaged-7nos, Vertical blue line-10nos, Frame bend-1no & blank display-1no & Frame insert problem 1no | 20 | Supplier planning to dispatch the replaced spares against invoice 25HGZ041501 on 23rd April'25
|  |  |  | 60.0005 | DOUBLE CHANNELS SLOTS-NEW TYPE | Connector damaged-11nos | 11
|  |  |  | 60.0006 | CONNECTOR PANEL Nellcor Ready | Sticker scratches-3nos | 3
|  |  |  | 60.0009 | Back Cabinet | Deep scratch mark-4nos, Cabinet top & right side buldge- 2 no's | 6
|  |  |  | 60.000547 | Rear connector | Connector damaged-7nos | 7
|  |  |  | 60.000549 | Arm board assembly with power source | Connector damaged-1no. Sequence number missing - 6 No's | 7
|  |  |  | 60.0005 | Cabinet assembly | Encoder damaged & Silicon keypad peeloff-3nos, silicon keypad black dots & Encoder damaged-4nos, Silicon keypad peeloff-1no., Cabinet black dot & encoder damaged-1no. Gap appear in top side - 1no's | 10
|  |  |  | 60.0009 | BACK CABINET | Deep scratch mark-4nos
2024-11-06 | 24ZX102101 | 50 | 1 | 60.0005 | LCD assembly | Vertical line - 1 No | 1
|  |  |  | 175.0001 | Fan for ultra Q5 | Fan Noise - 1 No | 1
|  |  |  | 60.000820 | IF board | Not getting on - 1 No | 1
|  |  |  | 60.0005 | Key board IF board | Speaker function not working. - 1 No | 1
|  |  |  | 60.0005 | ARM board | Not getting on - 1 No | 1
|  |  |  | 60.0005 | Alarm LED PCB | Oranage & green LED glowing together - 1 No | 1
|  |  |  | 60.0005 | Side key board | Key function not working properly - 1 No | 1

## M10 SKD issues
Received Date | BLT Invoice No. | Received Qty | Total SKD kits cannot be make saleable | P/N | Defective Item | Description | Failure Qty | Final Remarks
28-12-2024 | 24ZXS112101 | 500 | 85 | M1000104 | Main board without VGA & 3 in 1 interface | 1. Capacitor damage - 1No 2. Not switching ON - 1No 3. Connector damage( J-24,J-31, On/Off connector) - 3No's 4. USB port damaged(J-16) - 1No | 6 | Supplier planning to dispatch the replaced spares (Except LCD) against invoice 25HGZ041501 on 23rd April'25
|  |  |  | 1.14.101-0010-00-1 | 10.4' LCD | 1. Dull display - 74No's 2. Dot on the display.(Red, Green, Black) - 9No's 3. Rust on insert - 1No 4. Vertical line - 1No | 85
|  |  |  | M1000106 | SMPS module | Unit not swtching on | 1
|  |  |  | M1000108 | Side connector | Sticker Scratches - 21 nos | 21
|  |  |  | M1000103 | Back Cabinet with Battery | 1. Scratch on back cabinet -3No's 2. Black dot -1No | 4
|  |  |  | M1000102 | Front cover assembly without touchscreen | 1. Black dot on front frame - 2No's 2. Insert missing - 2No's 3. ON/OFF silica keypad scratch - 1No | 5
2024-10-26 | 24ZXS101501 | 50 | 28 | 1.14.101-0010-00-1 | 10.4' LCD | Brightness issue | 28
|  |  |  | 1.19.13-000001-0-1 | Front cover assembly with Touch Screen | touch option not working | 1
|  |  |  | 1.15.01-0063-01-00 | Power board (SMPS) | output not coming | 1

## Replacement Rejection
Model | Item No. | Description | Failure details | Qty | Import invoice & Date details | Final Remarks
TSIII -Upgradable | 60.0008 | BACK PANEL STICKER POWER SOURCE TS III UPGRADE | Power source assembly frame-TS III Back panel sticker scratches-3 nos. | 2 | SH241015001 dt. 14-10-2024 | Received
TSIII -Upgradable | 60.0008 | FRONT CABINET FOR TS III | Scratches-2nos, White dot-1no. | 3 | 24HGZ120601 dt. 06-12-2024 | Received
TSIII -Upgradable | 50.000260 | Backlight cable TS III | Shortage receipt | 2 | 24HGZ120601 dt. 06-12-2024 | Received
TSIII -Upgradable | 50.0003 | Battery lock of TS III | Wrong item recv'd | 4 | 24HGZ120601 dt. 06-12-2024 | Received wrongly instead of complete battery lock.
TSIII -Upgradable | 60.0008 | LCD assembly | Green dot-2nos. SN: NL-3U24100674, NL-3U24100701 | 2 | 24HGZ120601 dt. 06-12-2024 | Received
M10 | 1.14.101-0010-00-1 | 10.4' LCD | LCD Dull Display | 6 | 24HGZ121102 dt.11.12.2024 | Received
M10 | 1.19.13-000001-0-1 | Front cover & Speaker (without touchscreen) | Front cover checked and found to be ok. But speaker not ok due to not working. | 1 | 24HGZ121102 dt.11.12.2024 | Received
M10 | 1.14.101-0010-00-1 | 10.4' LCD | LCD Dull Display | 7 | 24HGZ121101 dt.11.12.2024 | Received
Mini | 60.0004 | DIGITAL SPO2 MODULE | Sensor not detected - Z20448737 | 1 | 24HGZ121102 dt.11.12.2024 | Received NIBP Module instead of Spo2 module
TSIII - Upgradable | 60.000840A | Front cabinet(cabinet assembly) | Cabinet Scratches | 2 | 24HGZ121102 dt.11.12.2024 | Received
TSIII - Upgradable | 50.0003 | Battery lock of TS III | Wrong item Rece'd | 5 | 24HGZ121102 dt.11.12.2024 | Received wrongly instead of complete battery lock.
TSIII - Upgradable | NA | ALARM LAMP COVER | Front cabinets replacement received without lamp cover | 27 sets | 24HGZ121102 dt.11.12.2024 | Received
TSIII - Upgradable | 60.000859A | Side Plate Assembly(2-IBP; 2-TEMP,analog spo2) | Side connector panel problem, I/F board & SPO2 Module and sticker scratches | 1 | 24HGZ121102 dt.11.12.2024 | Analog Interface board and analog module not received. Only sticker received.
TSIII - Upgradable | 175.0001 | SMPS Fan | Vibration problem | 1 | 24HGZ121102 dt.11.12.2024 | Received
TSIII - Upgradable | 50.000316A | BM-100A module(2-IBP,Nellcor) | module recv'd without shield. And also module not working | 1 | 24HGZ121102 dt.11.12.2024 | Received
TSIII - Upgradable | 50.000316A | BM-100A module(2-IBP,Nellcor) | module recv'd without shield. And also module not working | 1 | 24HGZ121102 dt.11.12.2024 | Received
TSIII - Upgradable | 50.000120 | SPO2 SOCKET - DOC -10 TYPE | Wrong item Blue colour connector digital type Received | 1 | 24HGZ121102 dt.11.12.2024 | Received
M10 | 1.14.101-0010-00-1 | 10.4' LCD with foam | LCD Dull Display | 1 | 24HGZ121102 dt.11.12.2024 | Received
TSIII - Non Upgradable | 60.000802A | LCD assembly with base plate | LCD base plate only recv'd, LCD Not received | 1 | 24HGZ121102 dt.11.12.2024 | Received
TSIII - Non Upgradable | 50.000260 | LVDS Cable | Wrong item 2nos recv'd | 2 | 24HGZ121102 dt.11.12.2024 | Received
TSIII - Non Upgradable | 60.000743A | MB350 arm board assembly-TS III lite | I/O board only recv'd. Arm board not recv'd | 1 | 24HGZ121102 dt.11.12.2024 | Received
TSIII - Non Upgradable | 60.000743A | MB350 arm board assembly-TS III lite | I/O board only recv'd. Arm board not recv'd | 1 | 24HGZ121102 dt.11.12.2024 | Received
TSIII - Non Upgradable | 60.000743A | MB350 arm board assembly-TS III lite | I/O board only recv'd. Arm board not recv'd | 1 | 24HGZ121102 dt.11.12.2024 | Received
Ultra Q5 (Digital) | 60.0005 | BOARD FOR DOUBLE CHANNELS SLOTS-NEW TYPE | Not Working & pin damage | 1 | 24HGZ121102 dt.11.12.2024 | Received
Ultra Q5 (Digital) | 60.0005 | Arm board assembly with power source | SMPS only recv'd. Arm board not received | 1 | 24HGZ121102 dt.11.12.2024 | Received
TSIII - Upgradable | 60.000854A | Back cabinet(cabinet assembly)-TS III | Cabinet Scratches | 2 | 24HGZ121102 dt.11.12.2024 | Received
TSIII - Upgradable | 175.0001 | Handle for TS III | Deep scratch mark visible in top side of cabinet(handle bottom) | 9 | 24ZXS121903 dt.19-12-2024 | Received